## Slide 1: Introduction to Fluorescence & Absorbance
Jacob Beal
iGEM Measurement Summer Webinar: Quantifying fluorescence and cell count with plate readers
July 14th, 2020
This document does not contain technology or technical data controlled under either U.S. International Traffic in Arms Regulation or U.S. Export Administration Regulations.

## Slide 2: Fluorescent Protein = Debugging Output
Up to ~15 boolean, ~3-4 precision quantification
[Tsien Lab]
[Weiss]
[Hasty]

## Slide 3: Picking your instrument
Of course, your resources and instructors matter most…
 | Fluorimeter | Plate Reader | Flow Cytometer | Fluorescence Microscope
Throughput | Low | High | High | Low - High
Resolution | Population | Population | Single Cell | Subcellular
Time Series | Yes | Yes | No | Yes
Dynamic Range | 2-3 logs | 2-3 logs | 3-6 logs | 1-2 logs
# cells | n/a | n/a | High | Low
Today’s focus

## Slide 4: How a plate reader works
microplate
excitation light / laser
movable carrier
absorbance
detector
fluorescence
detector
splitting
mirror
filters

## Slide 5: How Absorbance (OD) Works
Cells (and other particles) scatter and absorb light
  Optical Density (OD) = log10 (source / detector)
Fraction of light detected depends on particle density, particle opacity, path length
  Scattering also depends on frequency (typically 600nm)
Absorbance per cell linear at low OD (< ~0.5), highly non-linear at high OD (> ~2.0)
absorbance
detector
light source

## Slide 6: How Fluorescence Works
Great signal, proportional to number of molecules
Exact unit relation in context is highly sensitive, hard to calculate
fluorescent
molecule
light emission at lower frequencies
incoming excitation light

## Slide 7: Factors affecting fluorescence measures
Molecule:
  Excitation and emission spectrum
  Brightness / quantum yield
  Proper folding
  Quenching from binding to other molecules
Sample:
  pH, oxygenation
  Background fluorescence: media, cell (can vary with cell state)
  Other fluorescent molecules (spectral overlap)
  Sample volume, orientation
Instrument
  Excitation strength, frequency
  Emission filters, light path
  Overlap of excitation and emission
  Detector amplification
Fluorescence always needs calibration to independent standards & process controls

## Slide 8: Excitation and Emission Spectra
Excitation spectrum:
Relative efficacy in stimulating emission
Emission spectrum:
Distribution of stimulated light

## Slide 9: Excitation and Emission Spectra
Spectra should match your instrument capabilities, not overlap significantly
Minor overlap (recommend <3%) can be disentangled with a linear transform

## Slide 10: Poorly Separated Excitation & Emission
TagBFP and mCerulean can be distinguished, but are too close for good quantification

## Slide 11: Separated by Excitation
TagBFP and GFPmut3 are excited by different wavelengths, thus do not overlap in readings

## Slide 12: Separated by Emission
TagBFP and mAmetrine are excited similarly, but emission is sufficiently separated.
Useful if your instrument cannot do multiple independent excitations

## Slide 13: Recommended Fluorescent Proteins
Green Fluorescent Proteins
  BBa_E0040: GFPmut3 (Excit. 500 / Emiss. 513, brightness 35, maturation time 4.1 min, weak dimer).
  BBa_K864100: sYFP2 (Excit. 515 / Emiss. 527, brightness 68, maturation time 4.1 min).
  Calibrant: fluorescein
Red Fluorescent Proteins
  BBa_J06504: mCherry (Excit. 587 / Emiss. 610, brightness 16, maturation time 15 min, pKa 4.5).
  mKate2 (Excit. 588 / Emiss. 633, brightness 25, maturation time 20 min, pKa 5.4).
  If a slow maturation time is acceptable:
    BBa_E1010: mRFP1 (Excit. 584 / Emiss. 607, brightness 12.5, maturation time 60 min, pKa 4.5).
    mScarlet (Excit. 569 / Emiss. 594, brightness 70, maturation time 174 min, pKa 5.3).
  Calibrants: Texas Red, Nile Red
Blue Fluorescent Proteins
  BBa_K592100: TagBFP (Excit. 402 / Emiss. 457, brightness 33, maturation time 13 min, pKa 2.7).
  If a slow maturation time is acceptable:
    Cerulean3 (Excit. 433 / Emiss. 475, brightness 35, maturation time 70 min, pKa 3.2).
  Calibrants: Coumarin 30 (not yet verified)

## Slide 14: If you need other proteins:
For more proteins & spectra, fpbase.org is an excellent resource!

## Slide 15: Summary
Fluorescence is a valuable biological debugging tool
Different instruments are good for different purposes
  Plate readers are good for tracing time series
  Flow cytometers are good for quantifying cell behavior
  Microscopes are good for spatial arrangement and subcellular structure
Plate readers typically measure both absorbance (OD) and fluorescence
Fluorescence is affected by many factors, so measurements must be calibrated
Fluorescent proteins should be chosen to match calibrants and not interfere with each other

## Slide 16: Calibration of Plate Reader Fluorescence and OD Measurements
Jacob Beal
iGEM Measurement Summer Webinar: Quantifying fluorescence and cell count with plate readers
July 14th, 2020
This document does not contain technology or technical data controlled under either U.S. International Traffic in Arms Regulation or U.S. Export Administration Regulations.

## Slide 17: iGEM Plate Reader Calibration Protocol
Calibration with serial dilutions of cheap, stable materials
  GFP: fluorescein (Sigma 46970), RFP: Texas Red (Sigma S3388)
  OD: Monodisperse silica beads (Nanocym 950nm)
Produces MEFL units directly comparable w. flow cytometry, models
Highly replicable & debuggable measurements (1.8x geo.std.)
Validated with large-scale interlab study (244 institutions)
[Beal et al., ‘18, Beal et al., ‘19]
Notes: Previously we recommended monodisperse silica beads from Cospheric, but they no longer sell the right diameter.

## Slide 18: Calibration Process
Running the calibration protocol: https://2020.igem.org/Measurement/Protocols
  Serial dilution of monodisperse silica particles
  Serial dilution of each fluorescent dye
  Enter all values in provided Excel sheet to calculate
Instrument settings must be identical for calibration and experiment!
  Turn off instrument auto-calibration! (e.g., path length correction, auto-gain)
Re-run calibration at least monthly (preferably weekly), to ensure nothing changes

## Slide 19: Serial Dilutions
Example of a good fluorescence series
Stock concentrations:
OD: 3.00E+09 Particles / mL
Fluorescein: 10 μM
Texas Red: 2 μM
Should be adaptable to other cell types by changing particle diameter, other fluorescence by changing calibrating dye

## Slide 20: Example of Good Calibration
For most instruments, effective dynamic range of fluorescence is larger than for OD
3.26E+08 particles / Abs600
Valid for Abs600 > 0.043
4.70E+09 MEFL / a.u.
Valid for a.u. > 47

## Slide 21: Problem: Inconsistent Dilution
If not linear in log scale (except for saturation), check data entry and/or redo pipetting

## Slide 22: Problem: Narrow Range
Only about 4x from highest to lowest range  adjust instrument settings
  OD range should be at least 20x
  Fluorescence range should generally be larger, depending on machine

## Slide 23: Problem: Negative Values
Readings should all be positive  instrument needs maintenance / reconfiguration

## Slide 24: Automated Calibration Validation
https://github.com/iGEM-Measurement-Tools/Excel_Process_Validator
Found Excel file iGEM 2019 Plate Reader Fluorescence Calibration - Example.xlsx
All expected sheets are present
Template appears to be intact
All variables were extracted
Sufficient dynamic range of Abs600 calibration values: 31.46
Found a sufficiently long particle dilution slope from column 1 to 7
Computed mean particles / Abs600 is positive
All non-blank wells show significant cell counts
All validation checks passed for Abs600
Sufficient dynamic range of fluorescein calibration values: 4387.20
Found a sufficiently long fluorescein dilution slope from column 1 to 11
Computed mean MEFL / a.u. is positive
All validation checks passed for fluorescence

## Slide 25: What do the units mean?
Equivalent Particles
  Calibrated units are close to cell counts (but not exact)
  Measure can be increased by debris, packing, expression of opaque materials
“This sample is as opaque as a suspension of this many cell-like spheres”
E. coli ≠ silica spheres
debris
packing
opacity

## Slide 26: What do the units mean?
Molecules of Equivalent X (e.g., MEFL X = Fluorescein, METR X = Texas Red)
  Calibrated units are close to molecule counts (but not exact)
  Measure be decreased by fluorescence inhibitors (e.g., slow folding, lack of oxygen)
“This sample is as fluorescent as this many molecules of fluorescein”
Fluorescein ≠ GFP

## Slide 27: Summary
Plate reader calibration is simple, cheap, and reliable
  Use consistent settings and recalibrate at least monthly
Calibration can detect problems with instrument or configuration
  Automated validation software is provided by iGEM
Calibrated units are close to cell count & molecule count (but not identical)
  Some types of conditions can cause large value changes

## Slide 28: Interpreting & Debugging Plate Reader Data
Jacob Beal
iGEM Measurement Summer Webinar: Quantifying fluorescence and cell count with plate readers
July 14th, 2020
This document does not contain technology or technical data controlled under either U.S. International Traffic in Arms Regulation or U.S. Export Administration Regulations.

## Slide 29: Estimating cell and molecule counts
Raw readings include background, even after calibration
To get accurate estimates, subtract background:
Calibration Excel templates do cell estimate and MEFL conversion, but not WT subtraction

## Slide 30: Gene expression  geometric statistics
Use geometric statistics in analysis:
Strongest fraction may dominate population
Why geometric stats?
Complex catalytic reactions  multiply many rates:
Rexpress = R1R2R3R4R5…
Central Limit Theorem  converge to log-normal!
Gamma distribution bursting also implies geometric stats
[Davidsohn et al., 2015], [Beal, 2017]

## Slide 31: How to Compute Geometric Statistics
Geometric statistics are just normal (arithmetic) statistics on a logarithmic scale:
Geometric mean = 10^(mean(log10(data)))
Geometric std.dev. = 10^(std(log10(data)))
Consequences:
Error bars no longer “plus/minus”
Instead: k-fold “times/divide”
Example of geometric vs. arithmetic statistics on per-cell fluorescence data:

## Slide 32: What about zero?
Net absorbance, fluorescence may be at or below zero
  Invalid on the log scale!
Really this is just giving an instrument limit:
Interpret as “< value” rather than zero (e.g., “MEFL/cell < 470”, “cells < 1.3e6”)
% Values close to autofluorescence / media indistinguishable from background
% geometric, because cells dominate
autofluorescence_std = geostd(negative_control_replicates);
indistinguishable_MEFL = autofluorescence_mean*(autofluorescence_std^2-1);
% arithmetic, because instrument error dominates
media_std = std(media_replicates);
indistinguishable_cells = 2*media_std;

## Slide 33: Experimental vs. Process Controls
Experimental Controls:
Is my hypothesis true?
One control per factor under study
Best when new data
Control very close to experiment conditions
Process Controls:
Should I trust the data?
One control per assumption in study
Best when known value
Control should have minimal relation to experiment conditions

## Slide 34: Experimental vs. Process Controls
Experimental Controls:
Process Controls:
GFP
GFP
GFP
RFP
RFP
RFP
null transfection
wild type cells
ERF beads
media only
focusing fluid

## Slide 35: Example Data to Analyze
https://github.com/iGEM-Measurement-Tools/Fluorescence-Tutorials

## Slide 36: Sanity Check Control Values
Compare to calibrants to ensure instrument linear range
Compare positive to max number of proteins per cell:
  E. coli: 2e6
  Yeast: 6e7
  Human: 2e9
Negative control should be much smaller than positive
Problems with these values indicate likely process failure
Example: E. coli negative <1e3, positive >1e4
http://book.bionumbers.org/how-many-proteins-are-in-a-cell/

## Slide 37: Example of Experimental Data
Colonies have generally grown well
Constructs are covering full reasonable range of gene expression, except medium-low

## Slide 38: Summary
Calibrated units can be converted to estimates of cell / molecule count
Gene expression should be analyzed using geometric statistics
Low values are indistinguishable from background
Data can be interpreted in relation to realistic biological values
  Positive and negative controls can detect process failures
  Experimental values can be related to biological intuitions