## Slide 1: Introduction to Flow Cytometry
Jacob Beal
iGEM Measurement Summer Webinar: Quantifying fluorescence and cell phenotypes with flow cytometry
July 21st, 2020

## Slide 2: Reminder: picking your instrument
For more on fluorescence in general, see the first portion last week’s webinar
 | Fluorimeter | Plate Reader | Flow Cytometer | Fluorescence Microscope
Throughput | Low | High | High | Low - High
Resolution | Population | Population | Single Cell | Subcellular
Time Series | Yes | Yes | No | Yes
Dynamic Range | 2-3 logs | 2-3 logs | 3-6 logs | 1-2 logs
# cells | n/a | n/a | High | Low
Today’s focus

## Slide 3: Why flow cytometry? Distributions Matter!
Are these two groups of cells the same?
But they've got the same mean fluorescence…?

## Slide 4: Per-cell measurements reveal variation
Quantifying variation components requires large numbers of per-cell measurements.
Example: RNA Replicon Cotransfection
Example: Constitutive mKate in HEK293
Spectral Bleed
Autofluorescence
Dose/Resource Variation
Transfection
Efficiency
Cell Variation
Dose Variation
Quantization
Dose Efficiency
[Beal et al., 2012]
[Beal et al., 2015]

## Slide 5: Distributions also help quality control
By modeling expected distributions, can detect failure of batches or individual samples
Good Samples
Bad Samples
Distribution Model
Bad Transfection

## Slide 6: Reminder: Fluorescence Measurements
Great signal, proportional to number of molecules
Exact unit relation in context is highly sensitive, hard to calculate
  Molecule spectrum, brightness, folding, quenching
  Sample pH, oxygenation, cell autofluorescence
  Instrument excitation source, emission filters, detectors
Use proteins with good spectral separation
fluorescent
molecule
light emission at lower frequencies
incoming excitation light

## Slide 7: How Flow Cytometry Works
Instrument-specific challenges:
Lots of data points!
Individual cells behave (very) differently
“granularity”
fluorescence
particle size
Debris/pipetting contamination
Disruptive measurement

## Slide 8: How Flow Cytometry Works
[Image courtesy of Sauro Lab, UW]
laser path
dichroic mirror
detector
filter

## Slide 9: Many data points are not from cells!
probable debris
probable cells
“Gating” = filtering out debris, cell clumps, instrument contaminants, etc.

## Slide 10: Area, Width, and Height
As a cell passes the laser, detectors register a pulse “event”
Most instruments report area, width, and (max) height
  e.g., “FITC-A”, “FITC-W”, “FITC-H”
Area is proportional to total fluorescence
Width and height are flow-dependent, useful for gating doublets, debris, & clumps
Area
Time
Intensity
Width
Height

## Slide 11: Internal Instrument Gating
Most instruments allow internal gating, to record only some events
Mostly avoid this: better to filter out extra events than wish you hadn’t lost data
Use only a conservative forward scatter (particle size) trigger
  Too low  non-cell events overwhelm recording capability, cell events randomly dropped
  Too high  cell events dropped
Time
FSC-A Intensity
✓
✓
✓
trigger threshold

## Slide 12: Cell Size vs. Flow Cytometers
Most flow cytometers were designed for human cells (~10um)
Too small: bacteria aren’t always fully cleaned out
  Can add extra “wash wells” between experimental samples
Too large: muscle cells, neurons will clog the cytometer nozzle
  Can use nuclear localization and extract nuclei
  And remember to ask if your cells are allowed on the machine at all!
Well A6
Well A7
Well A8
Contamination
from A6!
[Data taken by Traci Haddock-Angelli in 2013]

## Slide 13: Summary
Flow cytometry produces large volumes of single-cell measurements
Distributions of single-cell behavior reveal model parameters, protocol problems
Flow cytometers measure pulse area, width, and height on multiple channels
  Forward scatter indicates particle area
  Flourescence is best measured with area
  Width and height are useful for gating
Beware well-to-well contamination with small cells (e.g., bacteria)

## Slide 14: Calibrating and Checking Flow Cytometry Data with TASBE Flow Analytics
Jacob Beal
iGEM Measurement Summer Webinar: Quantifying fluorescence and cell phenotypes with flow cytometry
July 21st, 2020

## Slide 15: Metrology vs. Flow Cytometry
Flow cytometry is great for per-cell measurements, but…
Arbitrary unit output depends on…
Instrument brand, configuration
Interference from other colors
Choice of instrument settings
Run-to-run calibration drift
Fortunately, this can be corrected…
Yellow (Arbitrary Units)
Blue (Arbitrary Units)
These are not
physical units!

## Slide 16: Calibrated Flow Cytometry
Calibration with standard materials:
  MEFL units from NIST-certified rainbow beads
  Background subtraction w. WT/NT
  Spectral compensation w. single positives
  Color comparison w. multi-color controls
Highly replicable measurements (~1.2x std.)
Multiple free & open software packages:
  TASBE Flow Analytics (BBN)
  CytoFlow (MIT)
  FlowCal (Rice)
[Roederer, ‘02; Wang et al., ‘08; NIST/ISAC, ‘12; Beal et al., ‘12; Castillo-Hair et al., ’16; Beal et al.,’19]
https://github.com/TASBE/TASBEFlowAnalytics

## Slide 17: Gating out non-cell / dubious events
Use your negative control for this
Usually better to throw away some cells than keep trash

## Slide 18: Compensating for Autofluorescence
Use your negative control for this

## Slide 19: Compensating for Spectral Overlap
Use a strong positive control for each color
Note: only linear when autofluorescence subtracted
[cf. Roederer 2002]

## Slide 20: Translating Other Fluorescence to MEFL
Only FITC channel (e.g. GFP) goes directly
Others obtained from triple/dual constitutive controls
Must have exact same constitutive promoter!
Must have a FITC control protein!
This is a new control that you need to add

## Slide 21: Fluorescent Beads  Absolute Units
Run beads every time: flow cytometers drift up to at least 20 percent!
Also can detect instrument problems, mistakes in settings
SpheroTech RCP-30-5A

## Slide 22: Result: Distributions in MEFL Units

## Slide 23: Cell Size Estimation
Forward scatter (FSC-A) can be calibrated to estimate diameter with sized beads
  Relationship is systematic, but nonlinear!

## Slide 24: Summary
Flow cytometry data needs to be calibrated in a multi-stage process:
  Gating to remove clumps and debris
  Subtract autofluorescence to remove background
  Linear transformation to remove spectral overlap
  Mapping colors to equivalent units (if desired)
  Conversion to equivalent fluorescent molecules via known-fluorescence beads
Variable-size beads can calibrate forward scatter to estimate cell diameter
Free software packages are available (e.g., TASBE Flow Analytics)

## Slide 25: Interpreting Flow Cytometry Data
Jacob Beal
iGEM Measurement Summer Webinar: Quantifying fluorescence and cell phenotypes with flow cytometry
July 21st, 2020

## Slide 26: Reminder: geometric statistics
Use geometric statistics in analysis:
Strongest fraction may dominate population
Why geometric stats?
Complex catalytic reactions  multiply many rates:
Rexpress = R1R2R3R4R5…
Central Limit Theorem  converge to log-normal!
Gamma distribution bursting also implies geometric stats
[Davidsohn et al., 2015], [Beal, 2017]

## Slide 27: Reminder: Computing Geometric Statistics
Geometric statistics are just normal (arithmetic) statistics on a logarithmic scale:
Geometric mean = 10^(mean(log10(data)))
Geometric std.dev. = 10^(std(log10(data)))
Consequences:
Error bars no longer “plus/minus”
Instead: k-fold “times/divide”
Example of geometric vs. arithmetic statistics on per-cell fluorescence data:

## Slide 28: Example Data to Analyze
https://github.com/iGEM-Measurement-Tools/Fluorescence-Tutorials

## Slide 29: Arabinose-TetR circuit: Sample Analyses
AraC
pBAD
TetR
pTet
pBAD
GFP
RFP
Ara

## Slide 30: Example Device Input/Output
Note: error bars show geometric standard deviation of per cell output, not across replicates

## Slide 31: Signal-to-Noise Ratio
How well can we distinguish signals?
Out - In  ΔSNR
ΔSNR  Potential Circuit Depth
“Off”
“On”
Signal
Noise
μg = geo. mean
σg = geo.std.dev.
[Beal, Frontiers, 2015]

## Slide 32: Example Device SNR
SNR = 5.7 dB
SNR = 0.0 dB
Interpretation: pBAD working well, TetR much less so.
TetR doesn’t seem to be turning on very strongly…

## Slide 33: Summary
Use geometric statistics to summarize flow cytometry data
Consider distribution of single-cell behaviors
Signal-to-noise ratio helps assess device performance
Calibrated measurements help suggest approaches for improvement